Danfysik System SYS8X00 Command Reference

1. Communication Interface

The control board uses standard serial interfaces RS-232C and RS-422. Communication is performed by transmitting ASCII characters, terminated by CARRIAGE RETURN (CR). Responses from the power supply are terminated by LINE FEED (LF) and CARRIAGE RETURN (CR), though a recent update notes that answers are now followed by 'cr' alone. An error message includes a "?BELL" (ASCII 7). None of the serial lines use hardware control signals.

Two data links are available:

- A REMOTE LINE, with either RS-232C or RS-422 communication (optically isolated).

- A local control panel line.

The active communication line is referred to as the "line-in-command."

2. Commands (Alphabetical Order)

AD – Read ADC Channel

- Function: Reads a value from an ADC channel.

- Syntax: AD X (where X is 0-9)

- Response: ch sp val lf cr

- Details:

  - Channel 0: Output current (I/In)*100

  - Channel 1: Tesla T*100

  - Channel 2: Output voltage (V/Vn)*100

  - Channel 3: Internal +15V supply (V*10)

  - Channel 4: Internal -15V supply (V*10)

  - Channel 5: Internal +5V supply (V*10)

  - Channel 6: Delta temperature (Deg.C*10) - includes sign (+/-)

  - Channel 7: Transistor bank Vce (V)

  - Channel 8: Optional Iout 16-bit (I/In)*100000

  - Channel 9: Aux Iout (control panel) (I/In)*120

- Note: If a value is prefixed with 'E', it indicates the last value read during an error condition. Read the channel again to get the actual value once the error is cleared.

ADR – Address (Read/Write)

- Function: Reads or sets the MPS address.

- Syntax (Read): ADR

- Syntax (Write): ADR XXX (Address: 00 to 63 decimal)

- Response (Read): address lf cr

- Details: Not valid in LALL mode. The first ADR command after LALL disables the LALL function. Writing an address selects that unit and deselects the previous one.

ASW – Answer Mode

- Function: Switches the remote line into auto-answer mode for certain setup commands.

- Syntax: ASW

- Details: (Remote line only). Suppressed in LALL mode. Commands that generate an answer in this mode: WA, PO+/-, W1, W2, W3.

BAUD – Set Baud Rate

- Function: Sets the baud rate on the remote line.

- Syntax: BAUD XXXXX

- Details: Valid rates: 19200, 9600, 4800, 2400, 1200, 600, 300, 150. An illegal setting will break communication.

CMD – Command Line

- Function: Returns the current control mode.

- Syntax: CMD

- Response: REM (Remote) or LOC (Local).

CMDSTATE – Command Line State

- Function: Returns a detailed current control state.

- Syntax: CMDSTATE

- Response: REMOTE, LOCAL, or LOCK.

ERRC – Error in Code

- Function: Sets controller to respond with coded error messages.

- Syntax: ERRC

- Details: (Remote line only). For use with a host computer that can decode error codes.

ERRT – Error in Text

- Function: Sets controller to respond with full-text error messages.

- Syntax: ERRT

- Details: (Remote line only). For use with simple terminals.

F – Off

- Function: Switches the main power contactor OFF.

- Syntax: F

- Details: Clears interlocks if the interlock condition is removed.

GOFF – Global Off

- Function: Switches off the main contactor after ensuring DAC output is at minimum.

- Syntax: GOFF

LALL – Listen All

- Function: Puts all connected controllers into a pseudo-addressed mode.

- Syntax: LALL

- Details: (Remote line only). All controllers respond to setup commands (except N). No answers are given. Disabled by the first subsequent ADR command.

LOC – Local Control

- Function: Switches the line-in-command to the local control panel.

- Syntax: LOC

LOCK – Lock

- Function: Locks the line-in-command to the local line.

- Syntax: LOCK

- Details: (Remote line only). Used to prevent remote access during local service. Should only be used in emergencies.

N – On

- Function: Switches the main power contactor ON (if no interlock errors are present).

- Syntax: N

- Details: Cannot be used in LALL mode. Usually requires a preceding F or RS command to clear latched errors.

NASW – No Answer Mode

- Function: Cancels the auto-answer mode (ASW).

- Syntax: NASW

- Details: (Remote line only).

NERR – No Error

- Function: Suppresses detailed error messages. Only "?BELL" is returned on error.

- Syntax: NERR

- Details: (Remote line only).

PO – Polarity Read

- Function: Reads the present output polarity status.

- Syntax: PO

- Response: + (Normal), - (Reverse), N or O (Neutral/Off).

PO+ / PO- – Polarity Change

- Function: Changes the output polarity (if the unit is equipped for polarity reversal).

- Syntax: PO+ or PO-

PRINT – Print

- Function: Returns two lines of internal unit information (e.g., serial number, type).

- Syntax: PRINT

- Details: (Remote line only). Service command.

RA – Read DAC Setting

- Function: Reads the register controlling the current setting DAC (16-bit or optional 18-bit).

- Syntax: RA

- Response: DDDDDD (6 digits)

- Details: Calculate output current: Io = In * (Y / 1,000,000), where Y is the read value.

RAR – Read Actual Ramp End

- Function: Reads the DAC end value set for ramping.

- Syntax: RAR

- Response: RAR 025352 (6 digits with leading zeros)

REM – Remote Control

- Function: Switches the line-in-command to the remote line.

- Syntax: REM

RLOCK – Remote Lock

- Function: Locks the line-in-command to the remote line, inhibiting the control panel from switching to local.

- Syntax: RLOCK

- Details: (Remote line only). Can only be switched off by a new line-in-command from the remote line (REM or LOC).

RR – Read Slew Rate

- Function: Reads the current ramp speed setting.

- Syntax: RR

- Response: RR 050 (3 digits, 0.1-10% per second, default 050=5%).

RS – Reset

- Function: Resets (clears) all latched interlocks.

- Syntax: RS

S1 – Status 1

- Function: Returns a 24-character status string of interlocks, polarity, and on/off status.

- Syntax: S1

- Response: String of '!' (error/non-normal) and '.' (normal) characters. Bit definitions are provided in the manual.

SYN – Synchronize

- Function: Synchronizes/initializes the serial communication controller.

- Syntax: ASCII code 16 (Ctrl+V)

- Details: Similar to an IEEE-488 IFC. Send, wait 150 µs, then send a CR to re-initiate communication.

TS – Trigger Sequence

- Function: Starts a programmed sequence from its first position.

- Syntax: TS stack (stack: 0-3)

UNLOCK – Unlock

- Function: Releases a LOCK condition set by the local control panel.

- Syntax: UNLOCK

- Details: (Remote line only). Should only be used in emergency situations to regain remote control.

VER – Version

- Function: Returns three lines of internal program information (version, copyright, date).

- Syntax: VER

- Details: (Remote line only). Service command.

WA – Write DAC Setting

- Function: Writes a value to the register controlling the current setting DAC.

- Syntax: WA XXXXXX (6 digits)

- Details: Value interpretation is the inverse of the RA command.

WAR – Write Actual Ramp End

- Function: Writes the DAC end value to be used by the ramp function.

- Syntax: WAR DACend (Value from 001000 to maximum current)

- Details: Use the MAX command to find the maximum allowed current.

WR – Write Slew Rate

- Function: Sets the ramp speed.

- Syntax: WR rampspeed (Value 001 to 100, representing 0.1% to 10% of rated current per second).

- Details: Default initial setting is 050 (5%).

?1 – Quick Read High-Definition ADC

- Function: Quick binary read of the high-definition current ADC (24-bit raw value).

- Syntax: ?1

?2 – Quick Read Status (Binary)

- Function: Quick binary read of the status (32-bit). LSB of first byte corresponds to first bit of S1.

- Syntax: ?2

?3 – Quick Read Combined (?1 + ?2)

- Function: Quick binary read combining status and high-def ADC (32-bit + 24-bit).

- Syntax: ?3

?4 – Read ADC 100% Calibration Value

- Function: Reads the stored raw ADC value at 100% output calibration point (24-bit).

- Syntax: ?4

- Details: Used to calculate actual current percentage: ((?1 / ?4) * 100). Use MAX command to convert to Amps.

3. Command Summary

| Command | Description | Command | Description |
| --- | --- | --- | --- |
| AD X | Read ADC channel (0-9) | N | Main Power ON |
| ADR | Read MPS address | NASW | No answer mode |
| ADR XXX | Write MPS address (00-63) | NERR | No error message |
| ASW | Enter answer mode | PO | Read polarity status |
| BAUD XXXXX | Set baud rate | PO+ | Change to normal polarity |
| CMD | Read current control mode (REM/LOC) | PO- | Change to reverse polarity |
| CMDSTATE | Read detailed control state | PRINT | Print unit info |
| ERRC | Coded error message | RA | Read DAC set value |
| ERRT | Text string error message | RAR | Read actual ramp end value |
| F | Main Power OFF | REM | Change to remote control |
| GOFF | Global OFF | RLOCK | Remote lock |
| LALL | Listen ALL (multi-unit) | RR | Read slew rate |
| LOC | Change to local control | RS | Reset interlocks |
| LOCK | Lock MPS in local control | S1 | Read status string |
| MAX | Read maximum allowed current | SYN | Synchronize communication (Ctrl+V) |
| ?1, ?2, ?3, ?4 | Quick binary read commands | TS stack | Trigger sequence |
|  |  | UNLOCK | Unlock MPS |
|  |  | VER | Read version info |
|  |  | WA XXXXXX | Write DAC set value |
|  |  | WAR XXXXXX | Write ramp end value |
|  |  | WR XXX | Write slew rate |

DANFYSIK A/S - DK-4040 JYLLINGE - DENMARK
Document: P80303ML / P80303WH

### 1. Power Control

| Command | Function | Syntax | Expected Answer / Notes |
| --- | --- | --- | --- |
| N | Main Power ON | N | No answer. Requires no interlock errors. Use after F or RS if errors were present. |
| F | Main Power OFF | F | No answer. Clears interlocks if the fault is gone. |
| GOFF | Global OFF (soft shutdown) | GOFF | No answer. Ensures output is at minimum before turning off. |
| RS | Reset interlocks | RS | No answer. Clears latched error states. |

### 2. Setting & Reading Current

| Command | Function | Syntax | Expected Answer / Notes |
| --- | --- | --- | --- |
| WA | Write current setpoint (DAC) | WA XXXXXX (6 digits) | No answer. Sets target current. Formula: Io = In * (value / 1,000,000). |
| RA | Read current setpoint (DAC) | RA | DDDDDD (6 digits). Convert back using the same formula. |
| WAR | Write ramp end value | WAR XXXXXX (≥001000) | No answer. Sets the final current for ramping. |
| RAR | Read ramp end value | RAR | RAR 025352 (6 digits with leading zeros). |
| WR | Write ramp speed (slew rate) | WR XXX (001-100) | No answer. Value = 0.1% to 10% of rated current per second. Default 050 (5%). |
| RR | Read ramp speed | RR | RR 050 (3 digits with leading zeros). |
| AD 0 | Read output current (relative) | AD 0 | DDD (0-999). Represents (I/In)*100. |
| AD 8 | Read high-resolution current (if equipped) | AD 8 | DDDDDD (0-999999). Represents (I/In)*100000. |
| ?1 | Quick binary read of high-def ADC (raw) | ?1 | 24-bit binary value. Use with ?4 for % calculation. |
| ?4 | Read ADC 100% calibration value (raw) | ?4 | 24-bit binary value. Max reference for ?1. |

### 3. Polarity Control (if equipped)

| Command | Function | Syntax | Expected Answer |
| --- | --- | --- | --- |
| PO | Read polarity | PO | + (Normal), - (Reverse), N or O (Neutral/Off). |
| PO+ | Set to Normal polarity | PO+ | No answer. |
| PO- | Set to Reverse polarity | PO- | No answer. |

### 4. Status & Monitoring

| Command | Function | Syntax | Expected Answer / Notes |
| --- | --- | --- | --- |
| S1 | Read status string (24 bits) | S1 | 24 chars of ! (error) and . (normal). Detailed interlock/polarity info. |
| ?2 | Quick binary status read | ?2 | 32-bit binary. Maps to S1 bits. |
| ?3 | Combined status + ADC read | ?3 | 32-bit status + 24-bit ADC (binary). Fast monitoring. |
| CMD | Read control mode | CMD | REM (Remote) or LOC (Local). |
| CMDSTATE | Detailed control state | CMDSTATE | REMOTE, LOCAL, or LOCK. |
| PRINT | Read unit info (serial/model) | PRINT | Two lines of text. |
| VER | Read firmware version | VER | Three lines of text (version, copyright, date). |

### 5. Error Handling

| Command | Function | Syntax | Expected Answer / Notes |
| --- | --- | --- | --- |
| ERRC | Enable coded error messages | ERRC | No answer. Afterwards, errors return a numeric code (1-16). |
| ERRT | Enable text error messages | ERRT | No answer. Afterwards, errors return a descriptive string. |
| NERR | Suppress detailed errors | NERR | No answer. Only ?BELL (ASCII 7) on error. |
| (Any error) | — | — | Response includes ?BELL. Check S1 for details. |

### 6. Addressing & Communication Control

| Command | Function | Syntax | Expected Answer / Notes |
| --- | --- | --- | --- |
| ADR | Read/Write unit address | ADR (read) / ADR XXX (write) | Read: address. Write: No answer. Range 00-63. |
| LALL | Listen All mode (multi-unit) | LALL | No answer. All units accept commands (except N). Disabled by first ADR. |
| ASW | Auto-answer mode for some commands | ASW | No answer. Enables answers for WA, PO+/-, W1-W3. |
| NASW | Disable auto-answer | NASW | No answer. |
| SYN | Synchronize serial line (reset) | Ctrl+V (ASCII 16) | Possible S. Wait 150 µs then send CR. |

### 7. Local/Remote & Locking

| Command | Function | Syntax | Expected Answer |
| --- | --- | --- | --- |
| LOC | Switch control to Local panel | LOC | No answer. |
| REM | Switch control to Remote line | REM | No answer. Fails if locked locally. |
| LOCK | Lock control to Local (from remote) | LOCK | No answer. Emergency use only. |
| UNLOCK | Unlock Local lock (from remote) | UNLOCK | No answer. Emergency use only. |
| RLOCK | Lock control to Remote | RLOCK | No answer. Prevents panel from taking control. |

## Complete Protocol for a Single Power Supply (Slave)

Here's the complete step-by-step protocol to fully control one specific power supply unit.

### 1. INITIALIZATION & SELECTION

plaintext

SYN                    # Ctrl+V (ASCII 16) - Optional: reset serial if needed

(Wait 150µs)

CR                     # Carriage return

REM                    # Switch to remote control mode

ADR XXX                # Select specific unit (XXX = 00-63)

Expected:

- REM (if you send CMD after)

- No response from ADR XXX (unless error)

### 2. READ CURRENT (in Amps)

#### Method A: Using RA command (DAC value)

plaintext

RA                     # Read current setpoint register

Response: DDDDDD (6 digits, e.g., 500000)

Calculation to Amps:

text

Actual Current (A) = Nominal Current × (RA_value ÷ 1,000,000)

Example: If nominal = 100A and RA returns 500000 → 100 × (500000/1000000) = 50A

#### Method B: Using AD 0 (relative reading)

plaintext

AD 0                   # Read channel 0 (output current)

Response: 0 DDD (3 digits, 000-999)

Calculation to Amps:

text

Actual Current (A) = Nominal Current × (AD0_value ÷ 100)

Example: If nominal = 100A and AD0 returns 750 → 100 × (750/100) = 750A? Wait, that's 750%!* *Note: AD0 shows (I/In)×100, so750` = 7.5×In. For In=100A, that's 750A. This may be unrealistic - check your unit's max rating.*

#### Method C: High-resolution (if equipped)

plaintext

AD 8                   # 16-bit ADC (if available)

Response: 8 DDDDDD (6 digits, 000000-999999)

Calculation:

text

Actual Current (A) = Nominal Current × (AD8_value ÷ 100,000)

### 3. WRITE/SET CURRENT (in Amps)

Step 1: Calculate DAC value

text

DAC_value = (Desired_Current ÷ Nominal_Current) × 1,000,000

*Round to nearest integer between 000001-999999*

Step 2: Send command

plaintext

WA XXXXXX             # Write DAC value (6 digits)

Example: For 35A on a 100A unit → (35÷100)×1000000 = 350000 → WA 350000

Step 3: Optional - Set ramp end point

plaintext

WAR XXXXXX            # Set ramp end value (same calculation as WA)

Step 4: Turn ON (if not already)

plaintext

N                     # Main power ON

### 4. READ/WRITE RAMP SPEED

#### Read current ramp speed:

plaintext

RR                    # Read slew rate

Response: RR DDD (3 digits with leading zeros, e.g., RR 050)

Interpretation:

text

Ramp Speed (% of In/sec) = RR_value ÷ 10

Example: RR 050 = 5.0% of nominal current per second

#### Write new ramp speed:

plaintext

WR DDD                # Write slew rate (001-100)

Example: For 2.5%/sec → 2.5×10 = 25 → WR 025

### 5. READ VOLTAGE

plaintext

AD 2                  # Read output voltage channel

Response: 2 DDD (3 digits, 000-999)

Calculation to Volts:

text

Actual Voltage (V) = Nominal Voltage × (AD2_value ÷ 100)

Example: If nominal = 50V and AD2 returns 800 → 50 × (800/100) = 400V

Note: Some units may show actual volts if switch #3 is set and special PROM installed.

### 6. READ STATE & FAULTS

#### Method A: Detailed status (S1)

plaintext

S1                    # Read 24-bit status string

Response: 24 characters of ! (error) and . (normal)

Key bits interpretation:

- Bit 1: Main Power OFF

- Bit 2: Polarity Normal

- Bit 3: Polarity Reversed

- Bit 11: DC Overcurrent

- Bit 16: Waterflow Failure

- Bit 20: Panic Button / Door Switch

- Bit 21: Magnet Waterflow Failure

- Bit 23: MPS Not Ready

#### Method B: Quick binary status

plaintext

?2                    # Binary status (32-bit)

Response: 4 bytes of binary data (maps to S1 bits)

#### Method C: Combined status + current

plaintext

?3                    # Status + ADC (fastest monitoring)

Response: 32-bit status + 24-bit ADC (binary)

### 7. READ/WRITE POLARITY

#### Read current polarity:

plaintext

PO                    # Read polarity status

Response: PO X where X = + (normal), - (reverse), N/O (neutral/off)

#### Change polarity (if equipped):

plaintext

PO+                   # Set to normal polarity

or

plaintext

PO-                   # Set to reverse polarity

Important: Ensure current is at zero before changing polarity.

### 8. ERROR HANDLING

#### Enable error messages:

plaintext

ERRT                  # Text error messages (recommended)

or

plaintext

ERRC                  # Coded error messages (1-16)

#### Check for errors after each command:

- Listen for ?BELL (ASCII 7) sound

- Check response for error codes

- Use S1 to identify specific faults

#### Clear errors:

plaintext

RS                    # Reset latched interlocks

F                     # Power OFF (also clears some errors)

N                     # Power ON again (if error cleared)

### 9. COMPLETE EXAMPLE SESSION

Controlling unit at address 23 (100A, 50V nominal):

plaintext

REM                    # Ensure remote control

ADR 23                 # Select unit 23

# Set slow ramp and current

WR 010                 # 1%/sec ramp speed

WA 250000              # Set to 25A (25% of 100A)

WAR 250000             # Same as end point (no ramp)

# Turn on

N

# Monitor

AD 0                   # Check current (should show ~250)

AD 2                   # Check voltage

S1                     # Check status

PO                     # Check polarity

# Change to 50A with ramp

WAR 500000             # Set new target to 50A

# Current will ramp from 25A to 50A at 1A/sec (1% of 100A/sec)

# Quick monitoring

?3                     # Fast status + current read

# When done

F                      # Power OFF

### 10. QUICK REFERENCE CARD

| Action | Command | Response | Calculation |
| --- | --- | --- | --- |
| Select unit | ADR XX | (none) | XX = 00-63 |
| Read current | RA | DDDDDD | A = In × (val/1e6) |
| Set current | WA DDDDDD | (none) | val = (A/In)×1e6 |
| Read voltage | AD 2 | 2 DDD | V = Vn × (val/100) |
| Read status | S1 | 24 chars | !=error, .=normal |
| Read polarity | PO | PO X | X=+, -, N, O |
| Set polarity | PO+ or PO- | (none) | (if equipped) |
| Read ramp | RR | RR DDD | %/sec = val/10 |
| Set ramp | WR DDD | (none) | val = (%/sec)×10 |
| Power ON | N | (none) | (if no errors) |
| Power OFF | F | (none) | Clears errors |
| Reset errors | RS | (none) | Clears latched |

### IMPORTANT NOTES:

- Nominal Values Required: You MUST know your unit's nominal current (In) and voltage (Vn) for accurate Amp/Volt calculations.

- Safety Interlocks: The N command will fail if any interlock is active. Check S1 first.

- Ramp Behavior: Current changes gradually at the set ramp speed. Use WA for immediate setpoint, WAR for ramped change.

- Binary Commands: ?1, ?2, ?3, ?4 are faster but return binary data requiring parsing.

- Emergency Stop: In case of emergency, use GOFF (soft shutdown) or hardware panic button.

- Lockout: If control panel is locked (LOCK), you need UNLOCK (emergency only) before remote commands.

This protocol gives you complete control over a single power supply unit for all typical operations.